## Slide 1: Preparing the optical calibration of the M4-Deformable mirror for the E-ELT
Briguglio, R., Pariani, G., Xompero, M., Riccardi, A.
ADONI meeting, Firenze 12-14 April 2016

## Slide 2: E-ELT & M4
M4: the E-ELT’s DM
  Shape: flat
  A few numbers:
  OD: 234cm, ID: 54cm
  # segments: 6
  # actuators 5316
  on a triangular pattern (3.15cm pitch, 52.5 cm projected on M1)
  Settling time< 1ms
  Fitt. Error, median seeing:
  145 nm WF
  Flattening spec:
  20 nm WF
  Pupil position midway M4/M5:
     Optical area moves on M4, depending on field
  M5: Tip-Tilt only

## Slide 3: E-ELT & M4
M4: the E-ELT’s DM
  Shape: flat
  A few numbers:
  OD: 234cm, ID: 54cm
  # segments: 6
  # actuators 5316
  on a triangular pattern (3.15cm pitch, 52.5 cm projected on M1)
  Settling time< 1ms
  Fitt. Error, median seeing:
  145 nm WF
  Flattening spec:
  20 nm WF
  Pupil position midway M4/M5:
     Optical area moves on M4, depending on field
  M5: Tip-Tilt only

## Slide 4: E-ELT & M4
M4: the E-ELT’s DM
  Shape: flat
  A few numbers:
  OD: 234cm, ID: 54cm
  # segments: 6
  # actuators: 5316
  on a triangular pattern (3.15cm pitch, 52.5 cm projected on M1)
  Settling time< 1ms
  Fitt. Error, median seeing:
  145 nm WF
  Flattening spec:
  20 nm WF
  Pupil position midway M4/M5:
     Optical area moves on M4, depending on field
  M5: Tip-Tilt only

## Slide 5: M4 project
Partners
  ADOPTICA as prime-contractor (companies consortium, MG+ADS)
  INAF as sub-contractor
Allocated Budget:
  ~30M€ (tot)
  ~ 1.3M€/7y (INAF)
  2.5FTE/y
INAF Team:
  Riccardi, Briguglio, Xompero (OAA)
  Pariani (OAB)
  0.5 more persons
  (to be hired)
INAF Tasks:
  AO expertise
  Error budget
  Optical design of Test Tower
  Test plan and procedures
  Support of Tower integration and characterization
  M4 optical test (2022)

## Slide 6: The INAF Team
Marco Xompero-OAA
Giorgio Pariani-OAB
Runa Briguglio-OAA
Armando Riccardi-OAA

## Slide 7: However, we fit well into an ADONI meeting
LBT Access road, february 2010

## Slide 8: Project timeline
Date | 08/10/2016 | 08/09/2017 | 08/10/2019 | 08/01/2020 | 08/04/2022
Deliverables | Optical Test Setup Critical Design Review | Final Design review | Optical Test Setup Review | Integration readiness review | In-Unit Optical Tests
Optical Test Setup Design | X |  |  |  | 
Optical test setup specification | X |  |  |  | 
Compliance Matrix | X |  | X |  | 
Engineering Budgets |  |  | X |  | 
Optical test setup Verification plan | X |  |  |  | 
Optical test setup engineering budgets | X |  |  |  | 
M4 Unit optical tests plan/procedure | X | X |  | X | 
Optical test |  |  |  |  | X

## Slide 9
Review of the test tower design
    Interferometer replaced
    Optical design: update/sensitivity/risks
    Opto-mechanical fine tuning
2. Test procedure
    Risks/unknowns (experience with LBT/VLT +M4 prototype)
    Simulations
Overview of on-going activity

## Slide 10: On-going activity 1 Test Tower Revision
lim =
$  ∞+

## Slide 11
TOP VIEW
M4
FM
3350
390
SIDE VIEW
Concept:
Newtonian telescope
F/3.3 beam
45° folding mirror
1.5 m parabolic collimator
Flat M4 to close the cavity
Test tower concept
Interferometer, F/3 beam
Parabola

## Slide 12: Test tower concept
Rotating-sliding assembly
M4
Parabola
Ref Flat mirror
Interferometer frame

## Slide 13: Test Tower Revision
Validation of Fizeau interf. proposed during Phase1
  Dynamic Fizeau: discarded
      Critical internal calibration procedure
      Calibration residuals too large for M4 REQ
  New solution: dynamic Twyman Green
Optical design updated
  Beam expander to implement TG interf.
Validation tests held @OABo & OAB
Thanks to E.Diolaiti, F.Cortecchia, M.Riva

## Slide 14: Fizeau?
Dynamic Fizeau interferometer discarded:
Reasons for exclusion: calibration
  Tip/tilt and focus entangled with calibration signal
  Coupling of calibration signal with current aberrations
  Time overhead and risks due to calibration
  (e.g. Mask management)
Comparison between PSI and DYN mode: tip-tilt and focus are shown up. Critical for flattening
Re-trace error (signature) measured through HO cavity: entangled with local WF. Critical for IF and flattening.
Detector mask affected by reconstruction issues, because of very large fringes density. Critical for mask management/schedule.

## Slide 15
Our background in DM optical calibration:
On-going activity 2 Test procedures definition
LBT (2x),
Mag
VLT-DSM
LATT
M4DP
1m diam,
672acts
@LBT, @Arcetri
1.2m diam,
1170 acts
@Garching
40cm diam,
19 acts
@Arcetri
2x 30cm x70cm
2x 111 acts
@Merate
Also tested by Merate
2009-2011
2013
2014
2015
Also:
TecnoINAF2010 was devoted to DM calibration
SPIE 2010
AO4ELT3- SPIE2014
ICSO2015
SPIE2016
AO4ELT4

## Slide 16
M4 is different: è petaloso!!!
  Segmented system (2x tested with M4DP)
      Management of block-diagonal command matrix
      Management of mask islands
      Measurement/correction of global shape with local commands
  Images composition (interf beam << M4)
      Measur./Correction of continuous modes on partial frame
      Management of CPA vs segment aberrations
  Very tight specifications
      20 nm WFE, including phasing+polishing
  Test tower: 6 m
      Convection, vibration, thermoelastic bending

## Slide 17
M4 is different:
  Segmented system (2x tested with M4DP)
      Management of block-diagonal command matrix
      Management of mask islands
      Measurement/correction of global shape with local commands
  Images composition (interf beam << M4)
      Measur./Correction of continuous modes on partial frame
      Management of CPA vs segment aberrations
  Very tight specifications
      20 nm WFE, including phasing+polishing
  Test tower: 6 m
      Convection, vibration, thermoelastic bending

## Slide 18
The optical calibration of M4DP was a lesson learned:
  Management of segments:
      Commands
      Masks
      InfFunct
  Measurement/unwrapping of diff piston
  Measurement of piston InfFunct
      LO Modes command amplitude must be << λ/4
      We need a new command basis
Optical setup @ Merate



OAA+OAB:
February-March 2015
Results
-Flattening: 12 nm RMS WFE
-Diff piston accuracy: 5 nm WF
-Test on convection noise
-Noise vs diff piston
-Optimal control of differential tilt:
InfFunct absolutely referenced in tilt with the not-active shell

## Slide 19
Results on the M4DP
1: Initial shape
2: 2x10 modes flattening
3: high orders flattening, no phasing
4: whole system flattened and co-phased
WFE: 12 nm RMS
70 cm

## Slide 20: Still many open points
How to measure effectively the piston IntMat, despite the interf. phase ambiguity?
How to compute global commands
from single local measurements?
What is the best management strategy for the interferometer island masks?
How to cope with bench noise?
Diff piston noise: 150 nm PtV@25Hz
Nλ jumps

## Slide 21
8s is a simulation SW producing phase maps of the OTT.
Written in IDL
Model from FEA and ZEMAX
The user may command/modify:
Tower geometry/ configuration
M4 actuator command
Optics position
Optics figuring and markers on/off
Vibration noise , diff piston on/off
Introducing 8s: Optical Test Tower Simulator
Optics position
OTT WFmap
ZEMAX/FE model
OTT geometry
Test procedures
Data processing
Simulation strategy
Noise
Act command
OTT Data
Segment mask
Green: input from operator
Notes: Espandere un pò le cose fatte...

## Slide 22
8s is a simulation SW producing phase maps of the OTT.
Written in IDL
Model from FEA and ZEMAX
The user may command/modify:
Tower geometry/ configuration
M4 actuator command
Optics position
Optics figuring and markers on/off
Vibration noise , diff piston on/off
Introducing 8s: Optical Test Tower Simulator
Phase map
Geometry
Interferogram
Act.Command
Simulator output
Notes: Espandere un pò le cose fatte...

## Slide 23
Introducing 8s: Optical Test Tower Simulator
8s is a simulation SW producing phase maps of the OTT.
Written in IDL
Model from FEA and ZEMAX
The user may command/modify:
Tower geometry/ configuration
M4 actuator command
Optics position
Optics figuring and markers on/off
Vibration noise, diff piston on/off
Notes: Espandere un pò le cose fatte...

## Slide 24
Introducing 8s: Optical Test Tower Simulator
8s is a simulation SW producing phase maps of the OTT.
Written in IDL
Model from FEA and ZEMAX
The user may command/modify:
Tower geometry/ configuration
M4 actuator command
Optics position
Optics figuring and markers on/off
Vibration noise, diff piston on/off
Notes: Espandere un pò le cose fatte...

## Slide 25
Alignment procedure with reference flat
Single segment Interaction Matrix
Single segment flattening
Tip/tilt IntMat with partial segment visibility
Flattening of differential tip/tilt with partial segment visibility
‘Piston friendly’ modified command matrix
Work in progress with 8s
Mmat = [v0 v1 v2]
Modified Mmat:
v0’ = recmat # piston
v1’ = v1-v0’
v2’ = v2
piston(v0)=1
piston(v1)=0
piston(v2)=0
Piece of cake, mostly a system validation
Notes: Espandere un pò le cose fatte...

## Slide 26
Alignment procedure with reference flat
Single segment Interaction Matrix
Single segment flattening
Tip/tilt IntMat with partial segment visibility
Flattening of differential tip/tilt with partial segment visibility
‘Piston friendly’ modified command matrix
Work in progress with 8s
TTmat = [(2x6)x5136]=
[Tip0 Tilt0 0 0 .....
0 0 Tip1 Tilt1 .....
0 0 0 0 .....
0 0 0 0 ..... ]
Tip = <Mmat|Tip> |Mmat>
Mmat = [v0 v1 v2]
Modified Mmat:
v0’ = recmat # piston
v1’ = v1-v0’
v2’ = v2
piston(v0)=1
piston(v1)=0
piston(v2)=0
Notes: Espandere un pò le cose fatte...

## Slide 27
Alignment procedure with reference flat
Single segment Interaction Matrix
Single segment flattening
Tip/tilt IntMat with partial segment visibility
Flattening of differential tip/tilt with partial segment visibility
‘Piston friendly’ modified command matrix: now to be tested with diff piston noise
Work in progress with 8s
Original command matrix:
Mmat = [v0 v1 v2]
V0: mode 0
V1: LO modes (1:40...)
V2 : others (41:end)
Modified command matrix:
|v0’> = <v1|Zpiston> | v1>
|v1’>= (1- <v0’|v1 > )v1>
|v2’>= |v2>
Mmat’ = [v0‘ v1 ‘ v2’]
<Zpiston|v0>= 1
<Zpiston|v1>= 0
<Zpiston|v2>= 0
Notes: Espandere un pò le cose fatte...

## Slide 28: What’s next?
microOTT for test cases assessment
  Reduced scale OTT or
  Test-specific laboratory setup
Study on system calibration with FEA
  Lead by C. Del Vecchio
  Well fitting M4 modularity

## Slide 29: Summary
INAF is involved in M4 project
    2.5 FTE, 1.3 M€ / 7 years
    Tasks: AO expertise, optical test
Current activity
    Review of the optical test tower
    Definition of the calibration procedures
Optical test of M4 is a challenge
    Segmented system, local visibility, global performance
    Tight specifications
    Simulation in progress to assess the strategy

## Slide 30: Thanks for your attention stay tuned for flattening in 2023

## Slide 31: Update of Parabolic mirror spec.
The tightest SPEC is SurfError:
Slope error and Curv Radius are met after SurfError is met
Two possible SPEC definitions
Via simulation:
  We define a polishing quality spatial distribution
  Montecarlo generation of phase maps
  Computation of subtraction residuals
  Check of specifications
Direct computation of subtraction residuals
S = PM Surf Map
dS= S- shift(S, 2pix, 2pix):
dS < 3 nm RMS
Pix:
resolution element
3 nm  6 nm WF
 12 nm WF DPass